Infrared No-Touch Forehead Thermometer performance test report

Sample name: Infrared No-Touch Forehead Thermometer

Model: PT9L

Date: July 11, 2024

Test Report Home Page

Page 1 of 5 pages

| Sample Name | Infrared No-Touch Forehead Thermometer | Sample No. | YP2407017 |
| --- | --- | --- | --- |
| Model | PT9L | sample quantity | 1 unit |
| Number | LOT No.：KE04 | Test Place | Area B,Testing Center |
| Environment | 16℃~24℃，40%RH~64%RH | Test Date | On July 10,2024, On July 11,2024 |
| Test Items | - Accuracy |  |  |
| Reference documents for the test | - Infrared No-Touch Forehead Thermometer Product Accelerated Testing Protocol |  |  |
| Test Conclusion | - The accuracy of the tested Infrared No-Touch Forehead Thermometer meets the requirements of the Infrared No-Touch Forehead Thermometer Accelerated Life Test Protocol. Report date: July 11,2024 |  |  |
| Remarks | The damp-heat aging test was carried out by Tianxu(Tianjin)Testing Technology Service Co.,LTD. |  |  |

Approver:                                        Prepared:

| 1 | Accuracy | 33.0℃±0.3℃； | 0.0℃ | Meet the requirement | — |
| --- | --- | --- | --- | --- | --- |
|  |  | 37.0℃±0.2℃； | -0.1℃ | Meet the requirement | — |
|  |  | 41.0℃±0.2℃ | -0.1℃ | Meet the requirement | — |
| Following of this page blank |  |  |  |  |  |

| The damp-heat Accelerated aging test Temperature: 40℃; Humidity: 95% RH; Time: 14.8h |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| 1 | Accuracy | 33.0℃±0.3℃； | -0.1℃ | Meet the requirement | — |
|  |  | 37.0℃±0.2℃； | -0.2℃ | Meet the requirement | — |
|  |  | 41.0℃±0.2℃ | 0.0℃ | Meet the requirement | — |
| Following of this page blank |  |  |  |  |  |

Photos of Samples and Accessories

Front of the Infrared No-Touch Forehead Thermometer Back of the Infrared No-Touch Forehead Thermometer

Photos of Samples and Accessories

Nameplate of the Infrared No-Touch Forehead Thermometer

Description of the Sample and Condition

- The tested sample includes: 1 pcs Infrared No-Touch Forehead Thermometer. - The basic functions of the sample are normal before and after the test.

Following of this page blank