## 包装类
潜在失效模式分析
(Design FMEA)
分类： | 包装类 |  |  | 编制人： | [8 empty] | FMEA编号： | PT9L-PDFM01
产品型号： | PT9L |  |  | 审核： | [8 empty] | FMEA版本： | V1.0
FMEA小组： | 丛明、侯广伟、万龙、刘钊
项目 | 功能 | 潜在失效模式 | 潜在失效结果分析 | 严 重 度 (S) | 分 类 | 失效机制/原因 | 频 度 (O) | 现行控制 |  | 探 测 度 (D) | R.P.N. | 拟采取措施 | 措施执行结果 |  |  |  | 措施是否执行 |  | 检查结果
[8 empty] | 预防 | 探测 |  |  |  | 严 重 度 (S) | 频 度 (O) | 探 测 度 (D) | R.P.N. |  | R.P.N.
彩包装 | 包装体温计 | 纸托架（含体温计）无法装入彩包装或装入彩包装后晃动量太大 | 生产人员无法包装 | 4 | C | 彩包装的尺寸设计不合理 | 2 |  | 实际装配 | 2 | 16
|  | 彩包装支撑强度不够 | 彩包装褶皱或破损 | 4 | C | 彩包装的材质强度不够 | 3 |  | 跌落试验、运输试验 | 2 | 24
|  | 标识不清楚 内容不全面 | 用户误用 | 9 | A | 说明书印刷油墨质量不好 未完全符合法规要求 | 4 |  | 目测检查 法规符合性检查 | 5 | 180 | 标识内容及可识别度检验 | 9 | 4 | 1 | 36 | 检查检验记录 |  | □
纸托架 | 装入体温计（电池）整机 | 体温计无法装入纸托架或装入纸托架后晃动量太大 | 生产人员无法包装 | 4 | C | 纸托架的尺寸设计不合理 | 2 |  | 实际装配 | 2 | 16
|  | 纸托架支撑强度不够 | 纸托架褶皱或破损 | 4 | C | 纸托架的材质强度不够 | 3 |  | 跌落试验、运输试验 | 2 | 24
铭牌 | 给用户关于机器的信息 | 铭牌在运输过程中脱落 | 用户无法识别体温计的相关法规信息 用户误用 | 4 | C | 铭牌背胶强度不够 | 3 |  | 跌落试验、运输试验 | 2 | 24
|  | 铭牌印刷内容容易被擦拭掉 |  | 4 | C | 铭牌未覆膜 | 2 |  | 依IEC60601-1 电气安全通用要求中7标记耐久性要求 对铭牌进行试验 | 2 | 16
|  | 标识不清楚 内容不全面 |  | 9 | A | 印刷油墨质量不好 未完全符合法规要求 | 4 |  | 目测检查 法规符合性检查 | 5 | 180 | 铭牌内容及可识别度检验 | 9 | 4 | 1 | 36 | 检查检验记录 |  | □
说明书 | 给用户关于机器使用的信息 | 说明书印刷不够清楚 | 用户无法识别体温计使用的相关信息 | 4 | C | 说明书印刷油墨质量不好 | 2 |  | 实际使用 | 1 | 8
|  | 标识不清楚 内容不全面 | 用户误用 | 9 | A | 说明书印刷油墨质量不好 未完全符合法规要求 | 4 |  | 目测检查 法规符合性检查 | 5 | 180 | 说明书内容及可识别度检验 | 9 | 4 | 1 | 36 | 检查检验记录 |  | □
隔板 | 装入包装箱 | 垫板无法装入包装箱或垫板装入包装箱后晃动量太大 | 生产人员无法包装 | 4 | C | 垫板尺寸设计不合理 | 2 |  | 实际装配 | 2 | 16
|  | 垫板支撑强度不够 | 垫板褶皱或破损 | 4 | C | 垫板的材质强度不够 | 2 |  | 跌落试验、运输试验 | 1 | 8
外箱 | 包装机器 | 包装箱无法装入外箱或包装箱装入外箱后晃动量太大 | 生产人员无法包装 | 4 | C | 外箱尺寸设计不合理 | 2 |  | 实际装配 | 2 | 16
|  | 外箱支撑强度不够 | 外箱褶皱或破损 | 4 | C | 外箱的材质强度不够 | 2 |  | 跌落试验、运输试验 | 1 | 8
标签 | 提供机器商品及合格的信息 | 标签在运输过程中脱落 | 用户无法识别相关信息 | 4 | C | 标签背胶强度不够 | 3 |  | 跌落试验、运输试验 | 2 | 24
|  | 标识不清楚 |  | 4 | C | 印刷油墨质量不好 | 2 |  | 实际使用 | 1 | 8
纸托板 | 支撑 | 支撑强度不够 | 纸托板褶皱或破损 | 4 | C | 纸托板的材质强度不够 | 2 |  | 跌落试验、运输试验 | 1 | 8
保护膜 | 保护外包装 | 保护膜破损 | 顾客满意度下降 | 4 | C | 原材不良 | 2 |  | 外观及功能检查 | 2 | 16
防拆签 | 粘合彩包装 | 封口签翘起 | 影响外观 | 4 | C | 标签背胶强度不够 | 2 |  | 实际使用 | 1 | 8
FMEA小组成员
成员 | 姓名 | 签名 | 日期 | [10 empty] | 措施是否执行的检查结果
硬件工程师 | 万龙 | [12 empty] | 检查人员签名
硬件主管 | 侯广伟 | [12 empty] | 审核人员签名
软件工程师 | 王迪
软件主管 | 侯广伟
结构工程师 | 刘钊
结构主管 | 侯广伟
Note on F6: Ranking 1 to 10. See "Severity" tab.
Note on G6: CC: Critical SC: Significant See "DFMEArefcard" in "ReadMe"
Note on I6: Ranking 1 to 10. See "Occurrence" tab.
Note on L6: Ranking 1 to 10. See "Design Detection" tab.
Note on M6: Risk Priority Number = Product of "Severity", "Occur" & "Detect*. RPN is used for determining priorities. High values indicate areas where measures should primarily be focused.
Note on J7: What is done to reduce the occurence of the failure or prevent it from happening
Note on K7: What is done to detect if the failure occurs

## 严重度(S)
后果 | 判定准则：后果的严重度 | 级别
无警告的严重危害 | 潜在失效模式影响体温计的安全运行和/或包含不符合法规情形，失效发生时无预警。会危及使用者。 | 10
有警告的严重危害 | 潜在失效模式影响体温计的安全运行和/或包含不符合法规情形，失效发生时有预警。会危及使用者。 | 9
很高 | 体温计无法运行，丧失基本功能，顾客很不满意。 | 8
高 | 体温计能运行，但性能下降。顾客很不满意。 | 7
中等 | 体温计能运行，但丧失部分功能。顾客不满意。 | 6
低 | 体温计能运行，但部分功能性能下降。顾客有些不满意。 | 5
很低 | 装配/外观或其他小细节不符合要求。多数顾客（超过75%）发现有缺陷 | 4
轻微 | 装配/外观或其他小细节不符合要求。多数顾客（大约50%）发现有缺陷 | 3
很轻微 | 装配/外观或其他小细节不符合要求。多数顾客（少于25%）发现有缺陷 | 2
无 | 无可识别的影响 | 1

## 频度(O)
失效发生的可能性 | 判定准则 | 可能的失效率 | PPM level | 级别
很高 | 持续性发生的失效 | ≥10% | 100000 | 10
|  | 5% | 50000 | 9
高 | 反复发生的失效 | 2% | 20000 | 8
|  | 1% | 10000 | 7
中等 | 偶尔发生的失效 | 0.5% | 2000 | 6
|  | 0.2% | 500 | 5
|  | 0.1% | 100 | 4
低 | 相对很少发生的失效 | 500ppm | 10 | 3
|  | 100ppm | 1 | 2
极低 | 失效不太可能发生 | ＜10ppm |  | 1
Note on D3: 00329: 历史上未使用过的技术或设计。
Note on D4: 00329: Failure is inevitable with new design, new application, or change in duty cycle/ operating conditions.
Note on D5: 00329: Failure is likely with new design, new application, or change in duty cycle/ operating conditions.
Note on D6: 00329: Failure is uncertain with new design, new application, or change in duty cycle/ operating conditions.
Note on D7: 00329: Frequent failures associated with similar designs or in design simulation and testing.
Note on D8: 00329: Occasional failures associated with similar designs or in design simulation and testing.
Note on D9: 00329: Isolated failures associated with similar designs or in design simulation and testing.
Note on D10: 00329: Only isolated failures associated with almost identical design or in design simulation and testing.
Note on D11: 00329: No observed failures associated with almost identical design or in design simulation and testing.
Note on D12: 00329: Failure is eliminated through preventative control.

## 探测度(D)
探测度 | 判定准则：被设计控制探测的可能性 | or % Chance | Criteria (Likelihood of Detection by Design Control) | Ranking
绝对不可能 | There is no Design Control -or- Control will not or cannot detect a potential cause/mechanism and subsequent Failure Mode. | <5% | There is no Design Control -or- Control will not or cannot detect a potential cause/mechanism and subsequent Failure Mode. | 10
极微小 | Very Remote chance the Design Control will detect a potential cause/mechanism and subsequent Failure Mode. | 15% | Design analysis/detection controls have a weak detection capability (not likely to detect at any stage); Virtual Analysis is not correlated to expected actual operating conditions. | 9
微小 | Remote chance the Design Control will detect a potential cause/mechanism and subsequent Failure Mode. | 25% | System/Product verification/validation after design freeze and prior to launch with pass/fail testing (Subsystem or system testing with acceptance criteria such as ride and handling, shipping, evaluation, etc.). | 8
很少 | Very Low chance the Design Control will detect a potential cause/mechanism and subsequent Failure Mode. | 35% | System/Product verification/validation after design freeze and prior to launch with test to failure testing (Subsystem or system testing until failure occurs, testing of system interactions, etc.). | 7
少 | Low chance the Design Control will detect a potential cause/mechanism and subsequent Failure Mode. | 45% | System/Product verification/validation after design freeze and prior to launch with degradation testing (Subsystem or system testing after durability test, e.g., function check). | 6
中等 | Moderate chance the Design Control will detect a potential cause/mechanism and subsequent Failure Mode. | 55% | System/Product validation (reliability testing, development or validation tests) prior to design freeze using pass/fail testing (e.g., acceptance criteria for performance, function checks, etc.). | 5
中上 | Moderate High chance the Design Control will detect a potential cause/mechanism and subsequent Failure Mode. | 65% | System/Product validation (reliability testing, development or validation tests) prior to design freeze using test to failure (e.g., until leaks, yields, cracks, etc.). | 4
高 | High chance the Design Control will detect a potential cause/mechanism and subsequent Failure Mode. | 75% | System/Product validation (reliability testing, development or validation tests) prior to design freeze using degradation testing (e.g., data trends, before/after values, etc.). | 3
很高 | Very High chance the Design Control will detect a potential cause/mechanism and subsequent Failure Mode. | 85% | Design and analysis/detection controls have a strong detection capability. Virtual analysis is highly correlated with actual or extended operating conditions prior to design freeze. | 2
几乎肯定 | Design Control will almost certainly detect a potential cause/mechanism and subsequent Failure Mode. | >95% | Failure cause or failure mode can not occur because it is fully prevented through design solutions (e.g., proven design standard, best practice or common material, etc.). | 1