## 总装类
潜在失效模式分析
(Design FMEA)
分类： | 总装类 |  |  | 编制人： | [8 empty] | FMEA编号： | PT9L-ADFM01
产品型号： | PT9L |  |  | 审核： | [8 empty] | FMEA版本： | V1.0
FMEA小组： | 丛明、侯广伟、万龙、刘钊
项目 | 功能 | 潜在失效模式 | 潜在失效结果分析 | 严 重 度 (S) | 分 类 | 失效机制/原因 | 频 度 (O) | 现行控制 |  | 探 测 度 (D) | R.P.N. | 拟采取措施 | 措施执行结果 |  |  |  | 措施是否执行 |  | 检查结果
[8 empty] | 预防 | 探测 |  |  |  | 严 重 度 (S) | 频 度 (O) | 探 测 度 (D) | R.P.N. |  | R.P.N.
下盖 | 固定及保护内部元件，支撑外形。 | 完整度被破坏，下盖碎裂 | 不能正常测量 | 9 | C | 1.壳体壁厚不足 2.使用材料强度不够 | 1 | 1.严控注塑工艺 2.注塑件外观检查 | 跌落试验及运输试验 | 2 | 18
|  | 外壳漏电流 | 患者、操作者触及到带电部件 | 9 | B | 外壳绝缘强度不够 | 3 | 1.严控注塑工艺 2.注塑件外观检查 | 电气安全试验 | 3 | 81 | 漏电流、电介质强度测试 | 9 | 3 | 1 | 27 | 检查产品检验标准的漏电流、电介质强度项目 | □ | □
|  | 上盖未固定住 | 上盖与下盖脱开 | 5 | C | 固定结构设计不合理 | 2 | 1.工程工艺控制 2.整机外观检查 | 跌落试验及运输试验 | 2 | 20
上盖 | 固定及保护内部元件，支撑外形。 | 完整度被破坏，上盖碎裂 | 不能正常测量 | 9 | C | 1.壳体壁厚不足 2.使用材料强度不够 | 1 | 1.严控注塑工艺 2.注塑件外观检查 | 跌落试验及运输试验 | 2 | 18
|  | 外壳漏电流 | 患者、操作者触及到带电部件 | 9 | B | 外壳绝缘强度不够 | 3 | 1.严控注塑工艺 2.注塑件外观检查 | 电气安全试验 | 3 | 81 | 漏电流、电介质强度测试 | 9 | 3 | 1 | 27 | 检查产品检验标准的漏电流、电介质强度项目 | □ | □
|  | 下盖未固定住 | 上盖与下盖脱开 | 5 | C | 固定结构设计不合理 | 2 | 1.工程工艺控制 3.整机外观检查 | 跌落试验及运输试验 | 2 | 20
电池盖 | 与下配合后固定电池 | 完整度被破坏，电池盖碎裂 | 无法遮挡电池 | 5 | C | 1.壳体壁厚不足 2.使用材料强度不够 | 1 | 1.严控注塑工艺 2.注塑件外观检查 | 跌落试验及运输试验 | 2 | 10
|  | 电池盖与下盖配合松，非常容易脱落 | 使用过程中电池盖脱落 | 8 | B | 电池盖与下盖卡钩配合尺寸设计不合理 | 3 | 1.工程工艺控制 2.整机外观检查 | 跌落试验及温湿试验 | 3 | 72 | 拉力测试 | 8 | 3 | 1 | 24 | 符合最低拉力要求 | □ | □
透镜 | 与上盖粘贴固定，支撑外形。印刷后优化显示效果 | 透镜断差大，易划手 | 顾客对外观不满 | 8 | B | 透镜与上盖粘贴开胶发生断差 | 5 | 粘贴后使用按压工装 | 1.拉力测试 2.外观检查 | 2 | 80 | 来料管控 | 8 | 5 | 1 | 30 | 符合最低拉力要求 | □ | □
|  | 印刷不良 显示效果差 | 顾客对外观不满 | 8 | C | 印刷不良 | 5 | 调整印刷工艺 | 外观检查 | 1 | 40
保护盖 | 保护探头 | 配合不紧 | 无法起到保护作用 | 8 | B | 卡口配合偏差 | 5 | 严控注塑工艺 | 拔出力测试 | 2 | 80 | 来料管控 | 8 | 5 | 1 | 30 | 符合最低拉力要求 | □ | □
开关帽 | 控制按键 | 按键动作不良 | 功能不良 | 9 | B | 装配不良 | 5 | 1.工程工艺控制 2.按键测试 | 1.跌落试验及运输试验 2.按键测试 | 2 | 90 | 对成品进行100%按键复测 | 9 | 5 | 1 | 45 | 符合产品品质标准 | □ | □
支架 | 固定传感器 | 传感器与支架同心度差 | 影响准确度 | 9 | B | 装配不良 | 5 | 1.工程工艺控制 2.黑体测试 | 1.跌落试验及运输试验 2.黑体测试 | 2 | 90 | 对成品进行100%黑体复测 | 9 | 5 | 1 | 45 | 符合黑体测量标准 | □ | □
压块 | 固定波导管 | 波导管固定偏差 | 影响准确度 | 9 | B | 装配不良 | 5 | 1.工程工艺控制 2.黑体测试 | 1.跌落试验及运输试验 2.黑体测试 | 2 | 90 | 对成品进行100%黑体复测 | 9 | 5 | 1 | 45 | 符合黑体测量标准 | □ | □
压块 | 固定传感器PCB | 传感器PCB固定不良 | 影响准确度 | 9 | B | 装配不良 | 5 | 1.工程工艺控制 2.黑体测试 | 1.跌落试验及运输试验 2.黑体测试 | 2 | 90 | 对成品进行100%黑体复测 | 9 | 5 | 1 | 45 | 符合黑体测量标准 | □ | □
正极弹簧 负极弹簧双联簧 | 连接电路 固定电池 | 1.弹簧脱落 2.电池容易脱落 | 不能正常测量 | 8 | C | 1.下盖无卡钩固定弹簧 2.下盖电池仓的长度 不符合设计要求 | 3 | 1.工程工艺控制 2.上电后性能测试 | 实际使用 | 2 | 48
自攻螺钉 | 固定PCB | 螺钉连接强度不够 | PCB跌落后散落 | 9 | C | 1.螺钉本身强度或螺纹强度不符合设计要求 | 1 | 1.工程工艺控制 | 跌落试验、运输试验 | 2 | 18
连接线 | 连接电路 | 连接线焊接后脱落或折断 | 不能正常测量 | 9 | C | 1.连接线焊接长度不符合设计要求 | 1 | 1.工程工艺控制 2.上电后性能测试 | 跌落试验、运输试验 | 2 | 18
保护膜 | 防尘进入机器 | 保护膜破损 | 顾客满意度下降 | 5 | C | 原材不良 | 3 | 工程工艺控制 | 外观及功能检查 | 3 | 45
FMEA小组成员
成员 | 姓名 | 签名 | 日期 | [10 empty] | 措施是否执行的检查结果
硬件工程师 | 万龙 | [12 empty] | 检查人员签名
硬件主管 | 侯广伟 | [12 empty] | 审核人员签名
结构工程师 | 刘钊
结构主管 | 侯广伟
Note on F6: Ranking 1 to 10. See "Severity" tab.
Note on G6: CC: Critical SC: Significant See "DFMEArefcard" in "ReadMe"
Note on I6: Ranking 1 to 10. See "Occurrence" tab.
Note on L6: Ranking 1 to 10. See "Design Detection" tab.
Note on M6: Risk Priority Number = Product of "Severity", "Occur" & "Detect*. RPN is used for determining priorities. High values indicate areas where measures should primarily be focused.
Note on J7: What is done to reduce the occurence of the failure or prevent it from happening
Note on K7: What is done to detect if the failure occurs
Note on T9: 00914: 出厂检验有“漏电流、电介质强度”
Note on V9: 00914: 出厂检验有“漏电流、电介质强度”
Note on T15: 00914: 出厂检验有“漏电流、电介质强度”
Note on V15: 00914: 出厂检验有“漏电流、电介质强度”
Note on T16: 00914: 出厂检验有“漏电流、电介质强度”
Note on V16: 00914: 出厂检验有“漏电流、电介质强度”
Note on T19: 00914: 出厂检验有“漏电流、电介质强度”
Note on V19: 00914: 出厂检验有“漏电流、电介质强度”
Note on T20: 00914: 出厂检验有“漏电流、电介质强度”
Note on V20: 00914: 出厂检验有“漏电流、电介质强度”
Note on T21: 00914: 出厂检验有“漏电流、电介质强度”
Note on V21: 00914: 出厂检验有“漏电流、电介质强度”
Note on T22: 00914: 出厂检验有“漏电流、电介质强度”
Note on V22: 00914: 出厂检验有“漏电流、电介质强度”

## 严重度(S)
后果 | 判定准则：后果的严重度 | 级别
无警告的严重危害 | 潜在失效模式影响血压计的安全运行和/或包含不符合法规情形，失效发生时无预警。会危及使用者。 | 10
有警告的严重危害 | 潜在失效模式影响血压计的安全运行和/或包含不符合法规情形，失效发生时有预警。会危及使用者。 | 9
很高 | 血压计无法运行，丧失基本功能，顾客很不满意。 | 8
高 | 血压计能运行，但性能下降。顾客很不满意。 | 7
中等 | 血压计能运行，但丧失部分功能。顾客不满意。 | 6
低 | 血压计能运行，但部分功能性能下降。顾客有些不满意。 | 5
很低 | 装配/外观或其他小细节不符合要求。多数顾客（超过75%）发现有缺陷 | 4
轻微 | 装配/外观或其他小细节不符合要求。多数顾客（大约50%）发现有缺陷 | 3
很轻微 | 装配/外观或其他小细节不符合要求。多数顾客（少于25%）发现有缺陷 | 2
无 | 无可识别的影响 | 1

## 频度(O)
失效发生的可能性 | 判定准则 | 可能的失效率 | PPM level | 级别
很高 | 持续性发生的失效 | ≥10% | 100000 | 10
|  | 5% | 50000 | 9
高 | 反复发生的失效 | 2% | 20000 | 8
|  | 1% | 10000 | 7
中等 | 偶尔发生的失效 | 0.5% | 2000 | 6
|  | 0.2% | 500 | 5
|  | 0.1% | 100 | 4
低 | 相对很少发生的失效 | 500ppm | 10 | 3
|  | 100ppm | 1 | 2
极低 | 失效不太可能发生 | ＜10ppm |  | 1
Note on D3: 00329: 历史上未使用过的技术或设计。
Note on D4: 00329: Failure is inevitable with new design, new application, or change in duty cycle/ operating conditions.
Note on D5: 00329: Failure is likely with new design, new application, or change in duty cycle/ operating conditions.
Note on D6: 00329: Failure is uncertain with new design, new application, or change in duty cycle/ operating conditions.
Note on D7: 00329: Frequent failures associated with similar designs or in design simulation and testing.
Note on D8: 00329: Occasional failures associated with similar designs or in design simulation and testing.
Note on D9: 00329: Isolated failures associated with similar designs or in design simulation and testing.
Note on D10: 00329: Only isolated failures associated with almost identical design or in design simulation and testing.
Note on D11: 00329: No observed failures associated with almost identical design or in design simulation and testing.
Note on D12: 00329: Failure is eliminated through preventative control.

## 探测度(D)
探测度 | 判定准则：被设计控制探测的可能性 | or % Chance | Criteria (Likelihood of Detection by Design Control) | Ranking
绝对不可能 | There is no Design Control -or- Control will not or cannot detect a potential cause/mechanism and subsequent Failure Mode. | <5% | There is no Design Control -or- Control will not or cannot detect a potential cause/mechanism and subsequent Failure Mode. | 10
极微小 | Very Remote chance the Design Control will detect a potential cause/mechanism and subsequent Failure Mode. | 15% | Design analysis/detection controls have a weak detection capability (not likely to detect at any stage); Virtual Analysis is not correlated to expected actual operating conditions. | 9
微小 | Remote chance the Design Control will detect a potential cause/mechanism and subsequent Failure Mode. | 25% | System/Product verification/validation after design freeze and prior to launch with pass/fail testing (Subsystem or system testing with acceptance criteria such as ride and handling, shipping, evaluation, etc.). | 8
很少 | Very Low chance the Design Control will detect a potential cause/mechanism and subsequent Failure Mode. | 35% | System/Product verification/validation after design freeze and prior to launch with test to failure testing (Subsystem or system testing until failure occurs, testing of system interactions, etc.). | 7
少 | Low chance the Design Control will detect a potential cause/mechanism and subsequent Failure Mode. | 45% | System/Product verification/validation after design freeze and prior to launch with degradation testing (Subsystem or system testing after durability test, e.g., function check). | 6
中等 | Moderate chance the Design Control will detect a potential cause/mechanism and subsequent Failure Mode. | 55% | System/Product validation (reliability testing, development or validation tests) prior to design freeze using pass/fail testing (e.g., acceptance criteria for performance, function checks, etc.). | 5
中上 | Moderate High chance the Design Control will detect a potential cause/mechanism and subsequent Failure Mode. | 65% | System/Product validation (reliability testing, development or validation tests) prior to design freeze using test to failure (e.g., until leaks, yields, cracks, etc.). | 4
高 | High chance the Design Control will detect a potential cause/mechanism and subsequent Failure Mode. | 75% | System/Product validation (reliability testing, development or validation tests) prior to design freeze using degradation testing (e.g., data trends, before/after values, etc.). | 3
很高 | Very High chance the Design Control will detect a potential cause/mechanism and subsequent Failure Mode. | 85% | Design and analysis/detection controls have a strong detection capability. Virtual analysis is highly correlated with actual or extended operating conditions prior to design freeze. | 2
几乎肯定 | Design Control will almost certainly detect a potential cause/mechanism and subsequent Failure Mode. | >95% | Failure cause or failure mode can not occur because it is fully prevented through design solutions (e.g., proven design standard, best practice or common material, etc.). | 1