Clinical accuracy report

iHealth PT9L Infrared No-Touch Forehead Thermometer

- Purpose

To verify the accuracy and repeatability of the Infrared No-Touch Forehead Thermometer, according to requirement in ATSEM 1965-98.

- Abbreviations

DUT: device under test, iHealth PT9L Infrared No-Touch Forehead Thermometer in this study

RCT: Reference clinical thermometer in this study

- Device tested description

2.1 Intended use

The Infrared No-Touch Forehead Thermometer is intended for the intermittent measurement of body temperature from the forehead skin surface on people of all ages. It can be used by consumers in the household environment and by healthcare providers.

2.2 Device description

The iHealth PT9L Infrared No-Touch Forehead Thermometer is hand-held, reusable, battery operated device, which can measure human body temperature on the skin temperature on one's forehead. Its operation is based on measuring the natural thermal radiation from the central forehead. Put the probe of the thermometer is onto the skin of a patient’s central forehead. Pressing the activation button to start the measurement of proposed infrared radiation. The electrical signal read out from the detector is fed to the circuit for amplification and calculation. The final measured temperature will be appeared on a LCD display.

The photo of the device is as follows:

2.3 Contraindications

It is not recommended for people whose forehead has local lesions, such as inflammation, trauma, recent surgery, etc..

2.4 See the user guide for how to operating the device.

- Reference clinical thermometer:

Electronic thermometer is used as reference clinical thermometer, using a water bath to verify the laboratory accuracy of the reference thermometer  before and after the test.

- Subject participate criteria

- A total number of subject is no less than 90.

- Within each age group both afebrile and febrile subjects were represented by nearly equal numbers.

- Oral temperature of 37.5℃ or higher is considered as febrile subjects.

- Subject group is as follows:

- Group 1 – infants, form 0-12 months old:  no less than 30 subjects

- Group 2 – children, form1-5 years old: no less than 30 subjects

- Group 3– adults, older than 5 years old: no less than 30 subjects

- Exclusive criteria

Subjects with medical conditions which might screw clinical accuracy validation results such as inflammation at the measuring site Subjects who have taken antipyretics in the preceding 120 min.

- Test Environment

Temperature is: 23±2℃

Humidity is : 55±10%

- Test method

Step 1: Sequence the Participants and sign Serial numbers for them according to participant time.

Step 2: using a water bath to verify the laboratory accuracy of the reference thermometer.

Step 3: The supervisor take the DUT, prepare to take measurement for the first subject.

Step 4: Perform three consecutive measurement procedures with the time interval no more than 1 min between each reading. And record the test results.

Step 5: Take one measurement with the RCT to get oral temperature, and record the test results.

Step 6: Repeat step 3 to step 5 to complete all the measurements of the participants.

Note, in this step, the participants signed with odd numbers were measured with DUT first and then with RCT, participants signed with even numbers were measured with RCT first and then with DUT.

Step 6.using a water bath to verify the laboratory accuracy of the reference thermometer again.

- Test results

  - Distribution of subjects

      - Group 1 – infants, form 0-12 months old

      - Group 2 – children, form1-5 years old

      - Group 3– adults, older than 5 years old

|  | Group 1 | Group 2 | Group 3 | Total |
| --- | --- | --- | --- | --- |
| Participants | 40 | 40 | 50 | 130 |
| F | 18 | 22 | 26 | 66(50.8%) |
| M | 22 | 18 | 24 | 64(49.2%) |
| Febrile | 16 | 15 | 20 | 51(39.2%) |
| Afebrile | 24 | 25 | 30 | 79(60.8%) |

  - Test results: There was no significant difference between the infrared thermometer and the oral temperature in different age groups. In the t test, p>0.05 means there is no significant difference between the two groups of data.

|  | Group 1 | Group 2 | Group 3 |
| --- | --- | --- | --- |
| Clinical bias(℃) | -0.07 | -0.05 | -0.05 |
| Standard deviation | 0.1 | 0.08 | 0.1 |
| Limits of agreement | 0.2 | 0.17 | 0.19 |
| Repeatability(℃) | 0.06 | 0.07 | 0.08 |
| P（t test） | 0.72>0.05 | 0.79>0.05 | 0.84>0.05 |
| Confidence interval of 95% | [-0.10,-0.03] | [-0.08,-0.03] | [-0.05,-0.02] |
| Clinical bias | -0.07 | -0.05 | -0.05 |
| Standard deviation | 0.1 | 0.08 | 0.1 |
| Limits of agreement | 0.2 | 0.17 | 0.19 |
| Uncertainty | +/-0.10 | +/-0.08 | +/-0.10 |

- Test data：In this test, Braun infrared thermometer NTF3000 was added.

| Group 1（Infants） |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| No. | Gender | Age | DUT |  |  | σj | σj2 | RCT | Braun | △cb1= ti1-RCT | △cb2= ti1-Braun | D1= ti2-ti1 | D2= ti3-ti2 | D3= ti1-ti3 |
|  | (F= Female/ M=Male) | (month) | ti1 | ti2 | ti3 |  |  | tRi |  |  |  |  |  |  |
| 1 | F | 3 | 38.2 | 38.3 | 38.2 | 0.058 | 0.003 | 38.3 | 38.4 | -0.10 | -0.20 | 0.1 | -0.1 | 0 |
| 2 | M | 2 | 36.8 | 36.6 | 36.6 | 0.115 | 0.013 | 36.8 | 36.9 | 0.00 | -0.10 | -0.2 | 0 | 0.2 |
| 3 | F | 1.5 | 36.8 | 36.9 | 36.9 | 0.058 | 0.003 | 36.9 | 37 | -0.10 | -0.20 | 0.1 | 0 | -0.1 |
| 4 | F | 1.5 | 38.7 | 38.7 | 38.5 | 0.115 | 0.013 | 38.7 | 38.8 | 0.00 | -0.10 | 0 | -0.2 | 0.2 |
| 5 | M | 2 | 38 | 38.2 | 38.2 | 0.115 | 0.013 | 38.2 | 38.3 | -0.20 | -0.30 | 0.2 | 0 | -0.2 |
| 6 | F | 1 | 37.8 | 37.8 | 37.7 | 0.058 | 0.003 | 37.7 | 37.8 | 0.10 | 0.00 | 0 | -0.1 | 0.1 |
| 7 | M | 0.5 | 38.8 | 38.8 | 38.7 | 0.058 | 0.003 | 38.8 | 38.9 | 0.00 | -0.10 | 0 | -0.1 | 0.1 |
| 8 | F | 0.5 | 38.5 | 38.6 | 38.6 | 0.058 | 0.003 | 38.7 | 38.6 | -0.20 | -0.10 | 0.1 | 0 | -0.1 |
| 9 | M | 3 | 36.8 | 36.9 | 36.8 | 0.058 | 0.003 | 37 | 37 | -0.20 | -0.20 | 0.1 | -0.1 | 0 |
| 10 | M | 1 | 37.6 | 37.5 | 37.5 | 0.058 | 0.003 | 37.5 | 37.7 | 0.10 | -0.10 | -0.1 | 0 | 0.1 |
| 11 | M | 1 | 36.6 | 36.7 | 36.7 | 0.058 | 0.003 | 36.8 | 36.8 | -0.20 | -0.20 | 0.1 | 0 | -0.1 |
| 12 | M | 2 | 36.7 | 36.8 | 36.7 | 0.058 | 0.003 | 36.8 | 36.9 | -0.10 | -0.20 | 0.1 | -0.1 | 0 |
| 13 | F | 2.5 | 36.5 | 36.5 | 36.6 | 0.058 | 0.003 | 36.5 | 36.6 | 0.00 | -0.10 | 0 | 0.1 | -0.1 |
| 14 | M | 0.5 | 36.9 | 36.9 | 37 | 0.058 | 0.003 | 37.1 | 37.2 | -0.20 | -0.30 | 0 | 0.1 | -0.1 |
| 15 | M | 0.5 | 37.2 | 37.3 | 37.2 | 0.058 | 0.003 | 37.1 | 37.3 | 0.10 | -0.10 | 0.1 | -0.1 | 0 |
| 16 | F | 1 | 36.8 | 36.8 | 36.9 | 0.058 | 0.003 | 36.9 | 37 | -0.10 | -0.20 | 0 | 0.1 | -0.1 |
| 17 | M | 1 | 36.6 | 36.6 | 36.6 | 0 | 0 | 36.7 | 36.8 | -0.10 | -0.20 | 0 | 0 | 0 |
| 18 | M | 3 | 36.6 | 36.6 | 36.7 | 0.058 | 0.003 | 36.6 | 36.8 | 0.00 | -0.20 | 0 | 0.1 | -0.1 |
| 19 | F | 2 | 37 | 37.1 | 37 | 0.058 | 0.003 | 37.1 | 37.1 | -0.10 | -0.10 | 0.1 | -0.1 | 0 |
| 20 | M | 0.5 | 38.3 | 38.2 | 38.2 | 0.058 | 0.003 | 38.3 | 38.4 | 0.00 | -0.10 | -0.1 | 0 | 0.1 |
| 21 | F | 7 | 37.9 | 38 | 38 | 0.058 | 0.003 | 38.1 | 38 | -0.20 | -0.10 | 0.1 | 0 | -0.1 |
| 22 | M | 8 | 38.5 | 38.5 | 38.4 | 0.058 | 0.003 | 38.5 | 38.6 | 0.00 | -0.10 | 0 | -0.1 | 0.1 |
| 23 | M | 10 | 38.8 | 38.9 | 38.9 | 0.058 | 0.003 | 38.9 | 39 | -0.10 | -0.20 | 0.1 | 0 | -0.1 |
| 24 | F | 10 | 38.7 | 38.7 | 38.5 | 0.115 | 0.013 | 38.6 | 38.8 | 0.10 | -0.10 | 0 | -0.2 | 0.2 |
| 25 | M | 10 | 37.8 | 37.8 | 37.7 | 0.058 | 0.003 | 37.8 | 37.9 | 0.00 | -0.10 | 0 | -0.1 | 0.1 |
| 26 | M | 10 | 36.8 | 36.8 | 36.7 | 0.058 | 0.003 | 36.8 | 36.9 | 0.00 | -0.10 | 0 | -0.1 | 0.1 |
| 27 | F | 5 | 36.5 | 36.6 | 36.6 | 0.058 | 0.003 | 36.7 | 36.6 | -0.20 | -0.10 | 0.1 | 0 | -0.1 |
| 28 | F | 5 | 36.3 | 36.4 | 36.3 | 0.058 | 0.003 | 36.5 | 36.5 | -0.20 | -0.20 | 0.1 | -0.1 | 0 |
| 29 | F | 9 | 36.9 | 36.8 | 36.8 | 0.058 | 0.003 | 36.8 | 37 | 0.10 | -0.10 | -0.1 | 0 | 0.1 |
| 30 | M | 10 | 36.2 | 36.3 | 36.3 | 0.058 | 0.003 | 36.2 | 36.3 | 0.00 | -0.10 | 0.1 | 0 | -0.1 |
| 31 | M | 10 | 36.7 | 36.8 | 36.7 | 0.058 | 0.003 | 36.8 | 36.9 | -0.10 | -0.20 | 0.1 | -0.1 | 0 |
| 32 | M | 10 | 36.9 | 36.9 | 37 | 0.058 | 0.003 | 37.1 | 37.1 | -0.20 | -0.20 | 0 | 0.1 | -0.1 |
| 33 | F | 10 | 36.3 | 36.4 | 36.3 | 0.058 | 0.003 | 36.4 | 36.5 | -0.10 | -0.20 | 0.1 | -0.1 | 0 |
| 34 | M | 8 | 36.6 | 36.6 | 36.6 | 0 | 0 | 36.7 | 36.8 | -0.10 | -0.20 | 0 | 0 | 0 |
| 35 | F | 6 | 36.1 | 36.1 | 36.2 | 0.058 | 0.003 | 36.1 | 36.2 | 0.00 | -0.10 | 0 | 0.1 | -0.1 |
| 36 | F | 5 | 37.6 | 37.7 | 37.7 | 0.058 | 0.003 | 37.6 | 37.7 | 0.00 | -0.10 | 0.1 | 0 | -0.1 |
| 37 | F | 10 | 37.8 | 37.9 | 37.8 | 0.058 | 0.003 | 37.9 | 38 | -0.10 | -0.20 | 0.1 | -0.1 | 0 |
| 38 | M | 9 | 38 | 38 | 38.1 | 0.058 | 0.003 | 38.1 | 38.2 | -0.10 | -0.20 | 0 | 0.1 | -0.1 |
| 39 | F | 10 | 36.5 | 36.4 | 36.5 | 0.058 | 0.003 | 36.6 | 36.7 | -0.10 | -0.20 | -0.1 | 0.1 | 0 |
| 40 | M | 10 | 36.8 | 36.8 | 36.9 | 0.058 | 0.003 | 36.8 | 36.9 | 0.00 | -0.10 | 0 | 0.1 | -0.1 |
| N=40 PT9L Repeatability: Sr = 0.06 |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| PT9L and RCT Clinical bias: Average(△cb1) = -0.07 Standard deviation:σ△cb1 = 0.10 Limits of agreement:LA1 = 0.20 Uncertainty: +/-0.10 |  |  |  |  |  |  | PT9L and Braun Clinical bias: Average(△cb2) = -0.15 Standard deviation:σ△cb2 = 0.06 Limits of agreement:LA2 = 0.13 Uncertainty: +/-0.06 |  |  |  |  |  |  |  |

| Group 2（Children） |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| No. | Gender | Age | DUT |  |  | σj | σj2 | RCT | Braun | △cb1= ti1-RCT | △cb2= ti1-Braun | D1= ti2-ti1 | D2= ti3-ti2 | D3= ti1-ti3 |
|  | (F= Female/ M=Male) | (year) | ti1 | ti2 | ti3 |  |  | tRi |  |  |  |  |  |  |
| 1 | M | 1.5 | 37.1 | 37.3 | 37.2 | 0.1 | 0.01 | 37.3 | 37.3 | -0.2 | -0.2 | 0.2 | -0.1 | -0.1 |
| 2 | F | 2 | 38.9 | 38.9 | 38.8 | 0.058 | 0.003 | 38.8 | 39 | 0.1 | -0.1 | 0 | -0.1 | 0.1 |
| 3 | F | 2 | 39.3 | 39.3 | 39.1 | 0.115 | 0.013 | 39.3 | 39.4 | 0 | -0.1 | 0 | -0.2 | 0.2 |
| 4 | M | 3 | 38.2 | 38.2 | 38.3 | 0.058 | 0.003 | 38.3 | 38.4 | -0.1 | -0.2 | 0 | 0.1 | -0.1 |
| 5 | M | 3 | 38.5 | 38.5 | 38.4 | 0.058 | 0.003 | 38.6 | 38.7 | -0.1 | -0.2 | 0 | -0.1 | 0.1 |
| 6 | F | 3 | 38.1 | 38 | 38.1 | 0.058 | 0.003 | 38.2 | 38.3 | -0.1 | -0.2 | -0.1 | 0.1 | 0 |
| 7 | M | 3 | 38.2 | 38.1 | 38.2 | 0.058 | 0.003 | 38.1 | 38.3 | 0.1 | -0.1 | -0.1 | 0.1 | 0 |
| 8 | F | 3 | 36.9 | 36.7 | 36.8 | 0.1 | 0.01 | 36.8 | 37 | 0.1 | -0.1 | -0.2 | 0.1 | 0.1 |
| 9 | F | 3 | 37 | 37 | 36.9 | 0.058 | 0.003 | 37 | 37.1 | 0 | -0.1 | 0 | -0.1 | 0.1 |
| 10 | F | 3 | 36.8 | 36.8 | 36.9 | 0.058 | 0.003 | 36.9 | 37 | -0.1 | -0.2 | 0 | 0.1 | -0.1 |
| 11 | M | 3 | 37.1 | 37.2 | 37 | 0.1 | 0.01 | 37.1 | 37.2 | 0 | -0.1 | 0.1 | -0.2 | 0.1 |
| 12 | F | 3 | 37 | 36.9 | 36.8 | 0.1 | 0.01 | 36.9 | 37.1 | 0.1 | -0.1 | -0.1 | -0.1 | 0.2 |
| 13 | F | 3 | 37 | 37.1 | 37.1 | 0.058 | 0.003 | 37.1 | 37.2 | -0.1 | -0.2 | 0.1 | 0 | -0.1 |
| 14 | F | 3 | 37.1 | 37 | 37 | 0.058 | 0.003 | 37 | 37.2 | 0.1 | -0.1 | -0.1 | 0 | 0.1 |
| 15 | F | 4 | 38.3 | 38.4 | 38.4 | 0.058 | 0.003 | 38.4 | 38.5 | -0.1 | -0.2 | 0.1 | 0 | -0.1 |
| 16 | M | 4 | 38.5 | 38.4 | 38.4 | 0.058 | 0.003 | 38.5 | 38.6 | 0 | -0.1 | -0.1 | 0 | 0.1 |
| 17 | M | 4 | 39 | 38.9 | 38.9 | 0.058 | 0.003 | 39 | 39.1 | 0 | -0.1 | -0.1 | 0 | 0.1 |
| 18 | F | 4 | 38 | 38.1 | 38 | 0.058 | 0.003 | 38 | 38.2 | 0 | -0.2 | 0.1 | -0.1 | 0 |
| 19 | F | 5 | 38.7 | 38.8 | 38.8 | 0.058 | 0.003 | 38.8 | 38.9 | -0.1 | -0.2 | 0.1 | 0 | -0.1 |
| 20 | F | 5 | 38.7 | 38.5 | 38.5 | 0.115 | 0.013 | 38.7 | 38.8 | 0 | -0.1 | -0.2 | 0 | 0.2 |
| 21 | F | 5 | 37 | 37 | 37 | 0 | 0 | 37 | 37.1 | 0 | -0.1 | 0 | 0 | 0 |
| 22 | M | 5 | 38.7 | 38.7 | 38.6 | 0.058 | 0.003 | 38.7 | 38.8 | 0 | -0.1 | 0 | -0.1 | 0.1 |
| 23 | F | 5 | 36.9 | 36.9 | 36.9 | 0 | 0 | 37 | 37.1 | -0.1 | -0.2 | 0 | 0 | 0 |
| 24 | M | 2.5 | 37.1 | 37 | 37.1 | 0.058 | 0.003 | 37.1 | 37.1 | 0 | 0 | -0.1 | 0.1 | 0 |
| 25 | M | 2.5 | 36.3 | 36.3 | 36.4 | 0.058 | 0.003 | 36.5 | 36.5 | -0.2 | -0.2 | 0 | 0.1 | -0.1 |
| 26 | F | 3.5 | 36.8 | 36.8 | 36.7 | 0.058 | 0.003 | 36.9 | 37 | -0.1 | -0.2 | 0 | -0.1 | 0.1 |
| 27 | M | 4.5 | 37 | 37 | 37.1 | 0.058 | 0.003 | 37.1 | 37.2 | -0.1 | -0.2 | 0 | 0.1 | -0.1 |
| 28 | F | 4.5 | 36.5 | 36.5 | 36.6 | 0.058 | 0.003 | 36.6 | 36.7 | -0.1 | -0.2 | 0 | 0.1 | -0.1 |
| 29 | M | 1.5 | 37 | 37.1 | 37.1 | 0.058 | 0.003 | 37.2 | 37.1 | -0.2 | -0.1 | 0.1 | 0 | -0.1 |
| 30 | F | 1.5 | 36.6 | 36.6 | 36.6 | 0 | 0 | 36.7 | 36.7 | -0.1 | -0.1 | 0 | 0 | 0 |
| 31 | F | 2 | 36.8 | 36.8 | 36.8 | 0 | 0 | 36.8 | 36.9 | 0 | -0.1 | 0 | 0 | 0 |
| 32 | M | 3 | 36.9 | 36.9 | 36.8 | 0.058 | 0.003 | 37 | 37 | -0.1 | -0.1 | 0 | -0.1 | 0.1 |
| 33 | M | 4 | 37.1 | 37.1 | 37 | 0.058 | 0.003 | 37.1 | 37.2 | 0 | -0.1 | 0 | -0.1 | 0.1 |
| 34 | M | 1.5 | 38 | 38 | 38.1 | 0.058 | 0.003 | 38.1 | 38 | -0.1 | 0 | 0 | 0.1 | -0.1 |
| 35 | F | 5 | 36.2 | 36.2 | 36.1 | 0.058 | 0.003 | 36.3 | 36.3 | -0.1 | -0.1 | 0 | -0.1 | 0.1 |
| 36 | F | 5 | 36.5 | 36.6 | 36.5 | 0.058 | 0.003 | 36.6 | 36.6 | -0.1 | -0.1 | 0.1 | -0.1 | 0 |
| 37 | M | 1.5 | 36.7 | 36.7 | 36.8 | 0.058 | 0.003 | 36.8 | 36.9 | -0.1 | -0.2 | 0 | 0.1 | -0.1 |
| 38 | M | 1.5 | 36.2 | 36.3 | 36.4 | 0.1 | 0.01 | 36.4 | 36.3 | -0.2 | -0.1 | 0.1 | 0.1 | -0.2 |
| 39 | F | 2 | 36.5 | 36.6 | 36.5 | 0.058 | 0.003 | 36.6 | 36.7 | -0.1 | -0.2 | 0.1 | -0.1 | 0 |
| 40 | M | 3 | 37.6 | 37.5 | 37.5 | 0.058 | 0.003 | 37.6 | 37.6 | 0 | 0 | -0.1 | 0 | 0.1 |
| N=40 Repeatability: Sr = 0.07 |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| PT9L and RCT Clinical bias: Average(△cb1) = -0.05 Standard deviation:σ△cb1 = 0.08 Limits of agreement:LA1 = 0.17 Uncertainty: +/-0.08 |  |  |  |  |  |  | PT9L and Braun Clinical bias: Average(△cb2) = -0.13 Standard deviation:σ△cb2 = 0.06 Limits of agreement:LA2 = 0.12 Uncertainty: +/-0.06 |  |  |  |  |  |  |  |

| Group 3（Adults） |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| No. | Gender | Age | DUT |  |  | σj | σj2 | RCT | Braun | △cb1= ti1-RCT | △cb2= ti1-Braun | D1= ti2-ti1 | D2= ti3-ti2 | D3= ti1-ti3 |
|  | (F= Female/ M=Male) | (year) | ti1 | ti2 | ti3 |  |  | tRi |  |  |  |  |  |  |
| 1 | M | 80 | 36.4 | 36.4 | 36.3 | 0.058 | 0.003 | 36.4 | 36.5 | 0.0 | -0.1 | 0 | -0.1 | 0.1 |
| 2 | F | 27 | 36.8 | 36.7 | 36.7 | 0.058 | 0.003 | 36.8 | 36.9 | 0.0 | -0.1 | -0.1 | 0 | 0.1 |
| 3 | F | 75 | 36.2 | 36.2 | 36.1 | 0.058 | 0.003 | 36.3 | 36.4 | -0.1 | -0.2 | 0 | -0.1 | 0.1 |
| 4 | F | 20 | 39.3 | 39.4 | 39.3 | 0.058 | 0.003 | 39.4 | 39.5 | -0.1 | -0.2 | 0.1 | -0.1 | 0 |
| 5 | M | 25 | 37.6 | 37.6 | 37.6 | 0 | 0 | 37.7 | 37.8 | -0.1 | -0.2 | 0 | 0 | 0 |
| 6 | F | 6 | 37.8 | 38 | 37.9 | 0.1 | 0.01 | 37.8 | 37.9 | 0.0 | -0.1 | 0.2 | -0.1 | -0.1 |
| 7 | M | 25 | 38.5 | 38.4 | 38.6 | 0.1 | 0.01 | 38.5 | 38.6 | 0.0 | -0.1 | -0.1 | 0.2 | -0.1 |
| 8 | M | 22 | 38.5 | 38.4 | 38.5 | 0.058 | 0.003 | 38.6 | 38.7 | -0.1 | -0.2 | -0.1 | 0.1 | 0 |
| 9 | F | 38 | 39.7 | 39.6 | 39.7 | 0.058 | 0.003 | 39.6 | 39.8 | 0.1 | -0.1 | -0.1 | 0.1 | 0 |
| 10 | M | 81 | 38 | 37.9 | 38 | 0.058 | 0.003 | 37.9 | 38.1 | 0.1 | -0.1 | -0.1 | 0.1 | 0 |
| 11 | F | 43 | 37.1 | 37.2 | 37.2 | 0.058 | 0.003 | 37.1 | 37.2 | 0.0 | -0.1 | 0.1 | 0 | -0.1 |
| 12 | F | 83 | 36 | 36 | 36.2 | 0.115 | 0.013 | 36.2 | 36.1 | -0.2 | -0.1 | 0 | 0.2 | -0.2 |
| 13 | F | 57 | 36.7 | 36.9 | 36.9 | 0.115 | 0.013 | 36.8 | 36.9 | -0.1 | -0.2 | 0.2 | 0 | -0.2 |
| 14 | F | 21 | 38.6 | 38.7 | 38.6 | 0.058 | 0.003 | 38.7 | 38.8 | -0.1 | -0.2 | 0.1 | -0.1 | 0 |
| 15 | M | 86 | 36 | 35.9 | 36.1 | 0.1 | 0.01 | 36.2 | 36.2 | -0.2 | -0.2 | -0.1 | 0.2 | -0.1 |
| 16 | M | 76 | 36 | 36.2 | 36.1 | 0.1 | 0.01 | 36 | 36.1 | 0.0 | -0.1 | 0.2 | -0.1 | -0.1 |
| 17 | F | 12 | 36.4 | 36.6 | 36.5 | 0.1 | 0.01 | 36.5 | 36.6 | -0.1 | -0.2 | 0.2 | -0.1 | -0.1 |
| 18 | M | 64 | 39 | 39.1 | 39.2 | 0.1 | 0.01 | 39 | 39.1 | 0.0 | -0.1 | 0.1 | 0.1 | -0.2 |
| 19 | F | 75 | 36.4 | 36.3 | 36.3 | 0.058 | 0.003 | 36.4 | 36.5 | 0.0 | -0.1 | -0.1 | 0 | 0.1 |
| 20 | F | 76 | 36 | 35.9 | 35.9 | 0.058 | 0.003 | 36 | 36.1 | 0.0 | -0.1 | -0.1 | 0 | 0.1 |
| 21 | F | 75 | 36.3 | 36.3 | 36.3 | 0 | 0 | 36.1 | 36.2 | 0.2 | 0.1 | 0 | 0 | 0 |
| 22 | M | 28 | 37.9 | 38 | 37.9 | 0.058 | 0.003 | 37.9 | 38 | 0.0 | -0.1 | 0.1 | -0.1 | 0 |
| 23 | F | 19 | 38.6 | 38.6 | 38.6 | 0 | 0 | 38.7 | 38.8 | -0.1 | -0.2 | 0 | 0 | 0 |
| 24 | M | 80 | 36.1 | 36.3 | 36.1 | 0.115 | 0.013 | 36.2 | 36.2 | -0.1 | -0.1 | 0.2 | -0.2 | 0 |
| 25 | M | 88 | 35.8 | 35.9 | 35.8 | 0.058 | 0.003 | 35.9 | 36 | -0.1 | -0.2 | 0.1 | -0.1 | 0 |
| 26 | M | 76 | 36.3 | 36.2 | 36.2 | 0.058 | 0.003 | 36.4 | 36.5 | -0.1 | -0.2 | -0.1 | 0 | 0.1 |
| 27 | F | 75 | 35.9 | 36 | 35.8 | 0.1 | 0.01 | 36 | 36.1 | -0.1 | -0.2 | 0.1 | -0.2 | 0.1 |
| 28 | F | 75 | 36.2 | 36.2 | 36.1 | 0.058 | 0.003 | 36.2 | 36.3 | 0.0 | -0.1 | 0 | -0.1 | 0.1 |
| 29 | M | 44 | 39.3 | 39.1 | 39.1 | 0.115 | 0.013 | 39.3 | 39.4 | 0.0 | -0.1 | -0.2 | 0 | 0.2 |
| 30 | M | 41 | 38.3 | 38.2 | 38.2 | 0.058 | 0.003 | 38.2 | 38.2 | 0.1 | 0.1 | -0.1 | 0 | 0.1 |
| 31 | F | 41 | 38.1 | 38 | 37.9 | 0.1 | 0.01 | 38.1 | 38.2 | 0.0 | -0.1 | -0.1 | -0.1 | 0.2 |
| 32 | F | 45 | 39 | 39 | 39 | 0 | 0 | 39 | 39.1 | 0.0 | -0.1 | 0 | 0 | 0 |
| 33 | F | 20 | 38.2 | 38 | 38.1 | 0.1 | 0.01 | 38.1 | 38.2 | 0.1 | 0.0 | -0.2 | 0.1 | 0.1 |
| 34 | M | 77 | 39.3 | 39.1 | 39.1 | 0.115 | 0.013 | 39.2 | 39.4 | 0.1 | -0.1 | -0.2 | 0 | 0.2 |
| 35 | M | 84 | 36.2 | 36.1 | 36.3 | 0.1 | 0.01 | 36.3 | 36.4 | -0.1 | -0.2 | -0.1 | 0.2 | -0.1 |
| 36 | F | 88 | 35.7 | 35.7 | 35.6 | 0.058 | 0.003 | 35.8 | 35.8 | -0.1 | -0.1 | 0 | -0.1 | 0.1 |
| 37 | F | 76 | 36.3 | 36.4 | 36.4 | 0.058 | 0.003 | 36.5 | 36.4 | -0.2 | -0.1 | 0.1 | 0 | -0.1 |
| 38 | M | 80 | 36 | 35.9 | 36 | 0.058 | 0.003 | 36.1 | 36.2 | -0.1 | -0.2 | -0.1 | 0.1 | 0 |
| 39 | F | 36 | 37.8 | 37.8 | 37.9 | 0.058 | 0.003 | 38 | 37.9 | -0.2 | -0.1 | 0 | 0.1 | -0.1 |
| 40 | M | 27 | 39.1 | 39.1 | 39.2 | 0.058 | 0.003 | 39.1 | 39.2 | 0.0 | -0.1 | 0 | 0.1 | -0.1 |
| 41 | F | 35 | 37.7 | 37.7 | 37.7 | 0 | 0 | 37.7 | 37.8 | 0.0 | -0.1 | 0 | 0 | 0 |
| 42 | F | 83 | 36 | 36 | 36.1 | 0.058 | 0.003 | 36.1 | 36.2 | -0.1 | -0.2 | 0 | 0.1 | -0.1 |
| 43 | M | 48 | 36.7 | 36.6 | 36.5 | 0.1 | 0.01 | 36.7 | 36.8 | 0.0 | -0.1 | -0.1 | -0.1 | 0.2 |
| 44 | M | 76 | 35.5 | 35.7 | 35.5 | 0.115 | 0.013 | 35.7 | 35.7 | -0.2 | -0.2 | 0.2 | -0.2 | 0 |
| 45 | M | 83 | 36.3 | 36.2 | 36.3 | 0.058 | 0.003 | 36.3 | 36.4 | 0.0 | -0.1 | -0.1 | 0.1 | 0 |
| 46 | M | 87 | 35.5 | 35.7 | 35.5 | 0.115 | 0.013 | 35.7 | 35.7 | -0.2 | -0.2 | 0.2 | -0.2 | 0 |
| 47 | M | 6 | 36.9 | 36.8 | 36.9 | 0.058 | 0.003 | 36.8 | 37 | 0.1 | -0.1 | -0.1 | 0.1 | 0 |
| 48 | F | 47 | 36.7 | 36.8 | 36.8 | 0.058 | 0.003 | 36.8 | 36.8 | -0.1 | -0.1 | 0.1 | 0 | -0.1 |
| 49 | F | 40 | 36.2 | 36.2 | 36.3 | 0.058 | 0.003 | 36.4 | 36.3 | -0.2 | -0.1 | 0 | 0.1 | -0.1 |
| 50 | M | 76 | 36 | 36 | 36.1 | 0.058 | 0.003 | 36.2 | 36.2 | -0.2 | -0.2 | 0 | 0.1 | -0.1 |
| N=50 Repeatability: Sr = 0.08 |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| PT9L and RCT Clinical bias: Average(△cb1) = -0.05 Standard deviation:σ△cb1 = 0.10 Limits of agreement:LA1 = 0.19 Uncertainty: +/-0.10 |  |  |  |  |  |  | PT9L and Braun Clinical bias: Average(△cb2) = -0.13 Standard deviation:σ△cb2 = 0.07 Limits of agreement:LA2 = 0.14 Uncertainty: +/-0.07 |  |  |  |  |  |  |  |