CHECKLIST of EN 60601-1-6:2010+A1:2015+A2:2021

Medical electrical equipment –

Part 1-6: General requirements for basic safety and essential performance –

Collateral standard: Usability

Product/project name：Infrared Forehead Thermometer

Model   name:         PT9L

Project number:      30000078

Conclusion(Pass/Fail):Pass

Drafted     by:                         Date              .

Reviewed by:                         Date              .

| No. | Version Date | Version Description | Version | Drafter | Approver |
| --- | --- | --- | --- | --- | --- |
| 1 | 2024.06.18 | Original | V1.0 | Wang Di | Hou Guangwei |

Version History

| IEC 60601-1-6 |  |  |  |
| --- | --- | --- | --- |
| Clause | Requirement + Test | Compliance method | Result |
| 4.0 | General requirements |  | P |
| 4.1 | Conditions for application to ME EQUIPMENT |  |  |
|  | The ME EQUIPMENT shall provide adequate USABLITY such that the RISks resulting from NORMAL USE, i.e. CORRECT USE and USE ERROR ,are acceptable. | Compliance with 4.2 and other clauses and subclauses of this collateral standard | P |
| 4.2 | USABILITY ENGINEERING PROCESS for ME EQUIPMENT |  |  |
|  | A USABILITY ENGINEERING PROCESS complying with IEC 62366-1:2015+A1:2020 shall be performed except: - the planning for and execution of production and POST-PRODUCTION monitoring in thecontext of applying the USABILITY ENGINEERING PROCESS within the framework ofISO 14971，and - maintenance of the USABILITY ENGINEERING PROCESS | USABILITY ENGINEERING PROCESS has been performed, please see the inspection of the USABILITY ENGINEERING file | P |
| 5 | ME EQUIPMENT ACCOMPANYING DOCUMENTS |  | P |
|  | The instructions for use include a brief description of the me equipment, its physical operating principles and significant physical and performance characteristics relevant to its use. | Inspect the instruction for use to include information required | P |
|  | The same information shall also be included in the technical description, if this is provided as a separate document. | Inspect the instruction for use to include information required | P |
|  | The instructions for use contain a summary of the USE SPECIFICATION as specified in IEC 62366-1:2015+A1:2020, 5.1. | Inspect the instruction for use to include use specification as specifed in IEC 62366-1. | P |

END OF EN 60601-1-6 REPORT.